## Slide 1: Plasma Diagnostic Techniques: Laser Induced Fluorescence (LIF/TALIF) spectroscopy
York Plasma Institute

## Slide 2
Laser Induced Fluorescence
dependent on:
- density
-
- volume
- collection optics

## Slide 3
Two-photon Absorption Laser-Induced Fluorescence (TALIF)
   Ground state density deduced by measuring the fluorescence signal and lifetime of the excited state
laser excitation 2 x hν
hν
Ground state
Excited state
De-excited state
TALIF: principle

## Slide 4
Quenching
hν
hν
Natural lifetime
Quenching coefficient
Quench partner density
Effective lifetime
Effective decay rate
Effective lifetime
Effective decay rate
Natural lifetime
Quench partner density
Quenching coefficient
TALIF
Collisional quenching
Collision
hν
Collisional quenching of laser excited state

## Slide 5
Example: Technological applications incl. biomedical
  Atoms: Precursors
Determination of N, O & H densities
Atmospheric pressure low-temperature plasmas

## Slide 6
RF μAPPJ
MP1101 BIOPLASMA
Parameters | Typical values
Excitation | 13.56 MHz
Power discharge | 1-2 W
Feed gas | 1 slm He + 0-1 % O2 – N2/O2
Channel size | 1 x 1 x 30 mm3
Gas temperature | » 340 K
Electron density | » 1017 m-3
Electron temperature | » 3 eV
6
electrons
ions
radicals
neutrals
ions
ions
ions
h
h
energy
radicals
metastables
neutrals
metastables
Core plasma
Plume
Liquid
Biological sample

## Slide 7
Plasma plume characterization:
Atomic densities
Oxygen
Synthetic air

## Slide 8
Picosecond-TALIF:
Experimental set-up
Photodiode
Variable
attenuator
μAPPJ
on 3-axis translation stage
Picosecond laser system:
30 ps pulses @ 10 Hz
λ: 193-2300 nm
Narrow bandwidth
Laser energy
feedback
8
ICCD detector, sub-nanosecond resolution
Interference filter
Pre-trigger
Delay generator
Example species
O 225.58 nm
N 206.65 nm
H 205.08 nm

## Slide 9
laser excitation 2 x 225.62 nm
844.87 nm
Ionisation threshold
Photo-ionisation
1 x 225.62 nm
TALIF signal
E2Laser
Picosecond-TALIF:
Oxygen TALIF scheme
2p4 3P2,1,0
3p 3P1,2,0
3s 3S
9
40 μJ
Notes: Signal prop to sq of energy/intensity – 2 photon
Energy is per time

## Slide 10
Picosecond TALIF: Signal acquisition
Camera gate width = 2 ns
Delay step = 1 ns
10
Laser shot

## Slide 11
Picosecond TALIF: Extraction of parameters
relative density = S x A*
11

## Slide 12
Picosecond-TALIF:
Calibration schemes
844.87 nm
2p4 3P2,1,0
3p 3P1,2,0
3s 3S
2 x 225.65 nm
Oxygen
834.91 nm
5p6 1S0
6p'[3/2]2
6s'[1/2]1
2 x 224.31 nm
Xenon
12
Excitation at the same wavelength
 important to keep the laser properties identical
K. Niemi et al., Plasma Sources Sci. Technol. 14 (2005) 375-386

## Slide 13
Picosecond-TALIF:
Calibration schemes
13
Excitation at the same wavelength
 important to keep the laser properties identical
742-746 nm
2p3 4S3/2
3p 4S3/2
3s 4P1/2,3/2,5/2
2 x 206.65 nm
Nitrogen
656.3 nm
1s 2S1/2
3d 2D3/2,5/2
2p 2P1/2,3/2
2 x 205.08 nm
Hydrogen
826.3 nm
4p6 1S0
5p’[3/2]2
5s’[1/2]1
2 x 204.13 nm
Krypton

## Slide 14
Picosecond-TALIF:
Calibration set-up
ICCD detector, sub-nanosecond resolution
Interference filter
834 nm
Pre-trigger
Delay generator
Photodiode
Variable
attenuator
Picosecond laser system:
30 ps pulses @ 10 Hz
λ: 224.28 nm
Laser energy
feedback
Calibration cell
filled with Xe
14

## Slide 15
Picosecond TALIF: Pressure inside the cell
Compromise between signal and lifetime
Use the same camera gate width for both O & Xe
15
pXe = 10 Torr

## Slide 16
Picosecond TALIF: Calibration with xenon
T: transmission of the detection optics
η: detector’s quantum coefficient
σ: two-photon absorption cross section
hν: incident photon energy
Aik: Einstein coefficient
τ: effective lifetime
S: fluorescence signal
16
0.5% oxygen
pXe = 10 Torr
From literature & calibration
Measured

## Slide 17: Measured vs calculated decay rate
  Uncertainty in the calculated decay rate
     Error introduced in the density calculation
 For reliable density measurement, decay rate should be measured
Difference between measured & calculated
  Too large to be due to uncertainty of He quenching coefficient
  does not depend on admixture
 Might be induced by 3 body collisions with He

## Slide 18: Measured vs calculated decay rate
  Uncertainty in the calculated decay rate
     Error introduced in the density calculation
 For reliable density measurement, decay rate should be measured

## Slide 19
Comparison:
Synchrotron VUV absorption
Synthetic air admixture variation
Quantitative and qualitative agreement
Synthetic air  complex gas mixture
19
Niemi et al. – 2013 – APL 103 034102

## Slide 20
Plasma plume characterization:
Oxygen vs synthetic air
Oxygen
Synthetic air

## Slide 21
Plasma plume characterization:
Oxygen vs synthetic air
Oxygen
Synthetic air

## Slide 22
Influence of surface on lifetime
No surface
Surface at 2.5 mm

## Slide 24
Thank you!